Smart Digital Stethoscope: Heart Sound Acquisition, Signal Processing and Quality Assessment

Senior Design Project. EEE 495 and EEE 496. Department of Electrical and Electronics Engineering.

Project Information

Course:  EEE 495 Senior Design Project I and EEE 496 Senior Design Project II. Two consecutive semesters, 28 teaching weeks in total.

Team size:  One to three students. See Section 6.

Tools:  C for the microcontroller firmware. Python (NumPy, SciPy, Matplotlib and a graphical user interface library) or MATLAB for the computer application. Bench instruments (oscilloscope, function generator, audio interface) are borrowed from the department. All software is free and open source, and no paid licence is required.

Data:  CirCor DigiScope Phonocardiogram Dataset on PhysioNet as the primary source. The PhysioNet or Computing in Cardiology 2016 heart sound recordings as an alternative. These recordings are open, anonymised and already published with their own ethical clearance. No data is collected from any person during this project.

Cost:  Approximately 100 to 200 US dollars in components, shared by the team. This covers the transducers, the analogue front end parts, a microcontroller development board, the prototyping boards, the phantom materials and an off the shelf stethoscope chestpiece. No instrument purchase is required.

Ethical position:  No recording is made from any human subject at any stage of this project. All physical validation is performed on an acoustic phantom that reproduces open access phonocardiogram recordings. No ethics committee application is therefore required. The prototype is a student device and not a certified medical device, it makes no diagnostic claim, and it is never applied to a person.

1. Topic

Auscultation remains one of the most widely used clinical examinations because it is non invasive, immediate and inexpensive. A digital stethoscope adds electronic amplification, recording, visualisation, digital filtering and objective quality assessment to that examination.

A digital stethoscope is not simply a microphone connected to a computer. Heart sounds occupy a low frequency band, roughly 20 to 500 Hz, at very low amplitude, and they are easily corrupted by contact noise, speech, movement and mains interference. The mechanical chestpiece, the acoustic coupling, the choice of transducer, the analogue front end and the sampling chain each change what is actually recorded. Recent work therefore treats transducer selection, diaphragm design and phantom based characterisation as separate practical design problems.

The engineering question of this project is whether a data acquisition chain built by a student team can capture heart sounds reliably enough to support reproducible measurement, and which design choices in that chain matter most. The answer is obtained by measurement on a repeatable acoustic phantom rather than by subjective listening.

2. Aim

To design, build and characterise a digital stethoscope and its complete acquisition chain, from the chestpiece to the computer application, and to determine by repeatable measurement which transducer, coupling and prototyping choices give the best signal quality within a stated cost and complexity budget.

3. Expected Outputs

The project produces the following, and is assessed on them:

- A working digital stethoscope prototype with acquisition electronics designed by the team, together with its embedded firmware and a documented device to computer protocol.

- A computer application providing live waveform and spectrogram display, recording with metadata, device control and recording quality feedback.

- A repeatable acoustic phantom that physically reproduces open access phonocardiogram recordings, together with the automated test scripts that use it.

- A characterisation report giving the measured frequency response, signal to noise ratio, repeatability, mains interference level, end to end latency and long term stability of the acquisition chain.

- A comparison of at least two transducer candidates and at least two prototyping approaches under identical phantom conditions, with the final design choice justified by measurement.

- A documented recording quality check that flags recordings unsuitable for further analysis, evaluated against controlled artefact injection.

- A complete source code repository, hardware documentation and reproducible setup instructions, and a final report.

Expected output at the end of Semester I:  A working acquisition chain from the phantom through the hardware and the microcontroller to the computer application, with a phonocardiogram captured at recognisable quality, a first characterisation of the chain, and an interim report.

Expected output at the end of the year:  The hardware hardened onto a soldered prototype board with the decision justified by measurement, the recording quality check evaluated, the full characterisation campaign completed including power and processing load, and a final report and defence.

4. Scope Boundaries

The following boundaries are fixed before the project starts. They answer the open questions raised in the original student proposal and they are not renegotiated during the year.

| Boundary | Decision |
| --- | --- |
| Cardiac sounds only | The project covers phonocardiography. Lung auscultation is outside the scope, because lung sounds occupy a different band with a different noise profile and would double the characterisation work. |
| No human subject recordings | All physical validation is performed on the acoustic phantom. No recording is made from any person, including members of the team. This removes the need for an ethics committee application and keeps the project schedule independent of any approval process. It also improves the quality of the evidence, because the phantom presents a known input that can be repeated exactly. |
| Electrical safety | The prototype is powered from a battery or from a computer USB port. It is not connected to the mains and it is not applied to a person, so no patient isolation barrier is required. The report states this limitation explicitly and discusses what would be needed for contact with a person, without claiming compliance with any medical device standard. |
| Photoplethysmography is optional | A synchronised optical pulse channel may be added in the second semester if the core device is complete and the schedule allows. It is not part of the assessed scope and its absence does not reduce the grade. |
| Machine learning is optional | A lightweight classifier may be trained and evaluated on the open dataset in the second semester if the schedule allows. It is not part of the assessed scope, it is evaluated on open data only, and no diagnostic claim is made. |
| A custom printed circuit board is not required | Development starts on a breadboard. The critical analogue section is transferred to a soldered prototype board in the second semester. A custom printed circuit board is optional. What is required is not the board itself but a documented decision, supported by noise floor, signal to noise ratio and repeatability measurements, on which level of construction is justified. |
| Mandatory measurements | Six quantities are reported for the final design: frequency response over 20 to 500 Hz on the phantom, signal to noise ratio, repeatability expressed as the spread over at least ten repeated measurements, mains interference level, end to end latency, and dropped packet rate over a continuous run. Classification performance is reported only if the optional machine learning module is attempted, and it is never a primary metric. |

5. Module Structure

| Module | Title | Responsibility |
| --- | --- | --- |
| A | Acquisition hardware and characterisation | Transducer comparison, the analogue front end, the mechanical coupling and chestpiece, construction of the acoustic phantom, bench measurement of noise floor, gain, bandwidth and repeatability, and the prototyping level decision. |
| B | Embedded acquisition and data path | Microcontroller sampling, buffering and timestamping, the device to computer protocol, device self test and status reporting, and measurement of latency, throughput and long term stability. |
| C | Computer application and quality assessment | Live waveform and spectrogram display, the recording library, playback of open dataset recordings through the phantom, the signal processing chain, the recording quality check, and the automated evaluation framework. |

6. Team Size and Module Allocation

The project is divided into three modules that are technically separable, so that the same plan works whether one, two or three students take the topic. Only the allocation changes.

- Three students: one module each. This is the intended configuration and gives the cleanest individual assessment.

- Two students: Module A is taken by one student and Module B by the other, and Module C is shared. Alternatively one student takes A and B while the other takes C.

- One student: the three modules are taken in sequence, the hardware is kept at breadboard and soldered board level only, and the second semester extensions are removed. The supervisor states this reduction in writing at the start of the project so that assessment remains fair.

7. Semester I, EEE 495: Weekly Schedule

| Week | Module A | Module B | Module C |
| --- | --- | --- | --- |
| MILESTONE (Week 1): team formed, modules allocated, project understood. |  |  |  |
| 1 | Read the proposal. List the transducer candidates and the bench equipment required. | Read the proposal. Identify the microcontroller options and the interface requirements. | Read the proposal. Identify the open datasets and the quality metrics to be used. |
| 2 | Study heart sound characteristics: amplitude range, frequency band and the sources of interference. | Study audio acquisition on a microcontroller: sampling, buffering and timestamping. | Download the dataset, write a reader and plot one recording. |
| 3 | Study electret condenser and micro electro mechanical microphones and select two candidates for comparison. | Study the communication options and draft the frame format for device to computer transfer. | Implement the quality metrics and verify them on recordings whose labels are known. |
| MILESTONE (Week 4): requirement specification frozen, components selected and ordered, tools installed by every member. |  |  |  |
| 4 | Write the requirements for the acquisition chain: gain, bandwidth, noise floor and clipping margin. | Write the requirements for the embedded layer: sampling rate, buffer depth, timestamp continuity and throughput. | Write the requirements for the computer application: display, recording, metadata and quality feedback. |
| MILESTONE (Week 5): bench setup operational, a known test signal captured from end to end. |  |  |  |
| 5 | Build the first breadboard front end. Measure the noise floor with the input shorted. | Bring up the development board. Stream a synthetic waveform to the computer and confirm that no samples are lost. | Build the live display and the recording function. Verify them with the synthetic stream from Module B. |
| 6 | Connect the first transducer. Measure gain and bandwidth with a function generator. | Replace the synthetic source with the real converter input. Verify timestamp continuity. | Add the spectrogram view and the recording library with metadata fields. |
| 7 | Build the acoustic phantom: a small transducer under a compliant coupling layer with the chestpiece above it. | Implement device status reporting and packet loss detection. | Implement the playback path so that a dataset recording can be driven into the phantom. |
| 8 | Measure the frequency response of the complete chain on the phantom using a swept signal. | Measure the end to end latency and run a continuous acquisition of at least thirty minutes. | Implement the automated test script that records, measures and writes the results table. |
| MILESTONE (Week 9): feasibility gate. A phonocardiogram played through the phantom must be captured with recognisable first and second heart sound structure. If this is not achieved, the transducer type or position is changed before any further work, and the scope is reduced if necessary. |  |  |  |
| 9 | Run the feasibility test. If it fails, change the transducer position or type before adding anything else. | Confirm that the recording is continuous, correctly timestamped and free of clipping. | Compare the captured waveform against the source recording and report the similarity numerically. |
| 10 | Compare the two transducer candidates on the phantom under identical conditions. | Add adjustable gain or filter modes if the selected chain supports them. | Implement the recording quality check and calibrate its thresholds on the phantom recordings. |
| 11 | Measure repeatability: repeat the same phantom measurement at least ten times and report the spread. | Test the behaviour under buffer stress and document how the device fails. | Add controlled artefact injection to the playback path: clipping, contact loss and interference. |
| 12 | Measure the mains interference level and the effect of grounding and cable routing. | Freeze the device to computer protocol and write its specification. | Evaluate the quality check against the injected artefacts and report its detection rate. |
| MILESTONE (Week 13): the baseline device works from end to end, from the phantom through the hardware and the microcontroller to the computer application, with a documented characterisation. |  |  |  |
| 13 | Write the hardware chapter of the interim report with the measured results. | Write the embedded chapter and the protocol specification. | Write the software and evaluation chapter with the first comparison tables. |
| 14 | Assemble the interim report; rehearse the design review presentation. | Assemble the interim report; rehearse the design review presentation. | Assemble the interim report; rehearse the design review presentation. |

8. Semester II, EEE 496: Weekly Schedule

| Week | Module A | Module B | Module C |
| --- | --- | --- | --- |
| MILESTONE (Week 1): Semester I reviewed, gaps identified, semester replanned. |  |  |  |
| 1 | Review the measured weaknesses of the front end and plan the move to a soldered board. | Review the embedded layer and list what needs improvement. | Review the quality check and the evaluation framework. |
| 2 | Transfer the critical analogue section to a soldered prototype board. | Improve buffering and error handling on the basis of the stress tests. | Extend the evaluation framework so that hardware revisions can be compared. |
| 3 | Repeat the noise floor, gain and bandwidth measurements on the soldered board. | Verify that the firmware behaves identically on the revised hardware. | Compare the two hardware revisions on the phantom under identical conditions. |
| 4 | Design and build the mechanical assembly: chestpiece, transducer mount and cable strain relief. | Add battery operation and measure its effect on the noise floor. | Extend the recording library with hardware revision metadata. |
| 5 | Measure the effect of coupling pressure and transducer position on the captured signal. | Implement the device self test: supply rails, saturation check and communication check. | Implement the ambient noise test: play the phonocardiogram through the phantom with controlled background noise. |
| 6 | Decide from the measurements whether a custom printed circuit board is justified, and document the decision. | Consolidate the firmware into a single documented build. | Consolidate the evaluation into a single script that produces every table in the report. |
| MILESTONE (Week 7): all subsystems implemented and individually verified. The characterisation campaign begins. |  |  |  |
| 7 | Freeze the hardware. Write its documentation and the bill of materials. | Freeze the firmware. Write its documentation. | Run the full characterisation sweep across transducers, revisions and noise conditions. |
| 8 | Support the campaign with any hardware configuration it requires. | Measure long term stability over a continuous run of several hours. | Produce the signal quality against ambient noise curves. |
| 9 | Measure power consumption and battery life. | Measure the processing load on the microcontroller and the remaining headroom. | Measure the processing time of the computer pipeline and confirm that it keeps up with the incoming data rate. |
| MILESTONE (Week 10): characterisation complete. Every claim in the report is supported by a measurement. |  |  |  |
| 10 | Write the hardware characterisation section. | Write the embedded performance section. | Write the system performance section. |
| 11 | Assist with the optional pulse channel if the schedule allows, otherwise support the report. | Implement the optional synchronised second channel if the schedule allows, otherwise support the report. | Evaluate the optional classifier on the open dataset if the schedule allows, otherwise support the report. |
| 12 | Prepare the demonstration setup and confirm that it is reliable outside the laboratory. | Prepare the demonstration firmware build and a fallback configuration. | Prepare the demonstration software mode and the offline playback fallback. |
| MILESTONE (Week 13): analysis complete. The design decisions are stated and justified by measurement. |  |  |  |
| 13 | Write the discussion of the transducer and front end decisions. | Write the discussion of the embedded design decisions. | Write the recommendation and the limitations section. |
| 14 | Finalise the report and the repository; rehearse the defence. | Finalise the report and the repository; rehearse the defence. | Finalise the report and the repository; rehearse the defence. |