Component and Laboratory Requirements

Smart Digital Stethoscope. Senior Design Project, EEE 495 and EEE 496.

Purpose of This List

This list expands Section 16 of the project proposal. The categories are unchanged. What is added here is the quantity of each item, an indicative cost, and the deadline by which it must be ordered.

The list is divided into three groups. Group A must be ordered by Week 4 of Semester I, because that week is a gate: a team that is still deciding then loses the semester to delivery times. Group B can wait until the second semester. Group C is equipment that is borrowed from the department and must not be purchased.

Costs are indicative only. They are given in US dollars for comparison and will vary with the supplier and with the exchange rate. The team should produce its own priced list before ordering.

Group A. To be ordered by Week 4 of Semester I

| Item | Purpose | Quantity | Cost, USD |
| --- | --- | --- | --- |
| Acoustic stethoscope | The chestpiece and tubing provide the mechanical coupling. A low cost item is sufficient; it will be dismantled. | 1 | 10 to 15 |
| Electret condenser microphone | Transducer candidate 1. Order two so that a damaged unit does not stop the work. | 2 | 10 to 60 |
| MEMS microphone, analogue or digital | Transducer candidate 2. Order two, for the same reason. | 2 | 10 |
| Low noise operational amplifier | The analogue front end. Order several; the gain is high and the first attempts are unlikely to survive. | 5 | 10 |
| Passive component kit | Resistors and capacitors for the front end, the bias network and the filters. | 1 set | 15 |
| Microcontroller development board | Embedded acquisition. A spare board is cheap insurance over a year long project. | 2 | 20 to 30 |
| Small transducer or exciter | The driver of the acoustic phantom. It reproduces the reference recordings physically. | 1 | 10 to 20 |
| Small audio amplifier module | Drives the phantom transducer. A class D module is sufficient. | 1 | 5 to 8 |
| Silicone sheet or gel pad | The compliant coupling layer of the phantom, between the driver and the chestpiece. | 1 set | 15 to 25 |
| Frame material for the phantom | A rigid frame that holds the chestpiece in a fixed position above the driver. Laser cut acrylic, 3D printed parts or plain wood are all acceptable. | 1 set | 10 |
| Breadboard and jumper set | First stage prototyping. | 1 set | 15 |
| Shielded cable and connectors | The connection between the transducer and the front end. Unshielded wire will pick up mains interference at this gain. | assorted | 10 to 15 |

Group B. Can wait until Semester II

| Item | Purpose | Quantity | Cost, USD |
| --- | --- | --- | --- |
| Soldered prototype board | Transfer of the verified analogue section from the breadboard. | 3 to 5 | 5 to 10 |
| Lithium ion cell with charge and protection module | Battery operation and its effect on the noise floor. | 1 set | 10 to 12 |
| Low noise voltage regulator | Supply rail for the analogue section. | 3 to 5 | 5 |
| Photoplethysmography module | Optional pulse channel. Outside the assessed scope; order only if the core device is complete. | 1 | 5 |
| Custom printed circuit board | Optional. Order only if the measurements justify it, as described in the scope boundaries. | optional | 20 to 40 |

Group C. May be borrowed from the department, not purchased

| Equipment | Used for |
| --- | --- |
| Oscilloscope | Noise floor, supply ripple and timing measurements. |
| Function generator | Gain and frequency response measurement of the acquisition chain. |
| Multimeter | Supply voltages and continuity. |
| Soldering station and solder | Assembly of the prototype board. |
| USB audio interface | Reference measurement of the transducers. Useful but not essential. |

What This List Adds to the Original Proposal

The categories in the proposal are complete. Five details were implicit and are made explicit here, because each of them has caused delay in comparable projects.

•  Passive component kit as a separate line. The proposal names the operational amplifier but not the resistors and capacitors around it.

•  An amplifier module for the phantom. The proposal specifies a small transducer for the phantom but not the means of driving it at a controlled level.

•  A rigid frame for the phantom. A phantom that is reassembled by hand between measurements cannot give a repeatable result, and repeatability is one of the six mandatory measurements.

•  Shielded cable between the transducer and the front end. At the gain this project requires, unshielded wire will collect mains interference that no amount of filtering will fully remove.

•  Spare units of the microphones, the operational amplifiers and the development board. These are inexpensive and a damaged part otherwise stops the work for as long as delivery takes.

Notes Before Ordering

•  Low frequency response is the selection criterion for the microphones. Heart sounds lie roughly between 20 and 150 Hz, while most commercial microphones are designed for speech and attenuate below 100 Hz. Before ordering, extract the lower minus 3 decibel point of each candidate from its datasheet and record it. A microphone that rolls off at 100 Hz will discard the strongest part of the signal, and the cause will not be obvious when the feasibility test fails in Week 9.

•  Order two of everything that carries signal. The front end runs at high gain and the first prototypes are not expected to survive.

•  Check what the department already holds before ordering. A development board, a breadboard set or a small speaker may already be available.

•  Keep every receipt and record the supplier and the part number. The bill of materials is a deliverable of the final report.

Cost Summary

Group A comes to approximately 150 to 200 US dollars. Group B adds approximately 30 to 70 US dollars, depending on whether the optional items are taken. The total for the year is therefore approximately 180 to 270 US dollars, or roughly 60 to 90 US dollars per student in a team of three.

Group C may be borrowed and costs nothing. If any item in Group C is not available in the department, the team must report this in the first three weeks so that an alternative can be arranged.